Persistence of Indentation with Bioabsorbable Poly-L/D-Lactide vs. Silicone Sponge Scleral Buckling Implants

Satu Länsman1, MD, Ari Karttunen2, MD, PhD, Heli Hirvelä1, MD, PhD, Tapani Palosaari1, MD, Minna Kellomäki3, Dr.Tech, Ville Ellä3, M.Sc., Pasi Ohtonen4 MSc, Pertti Törmälä3, PhD, MDhc, Timo Waris4 MD, PhD, Nureddin Ashammakhi3,4, MD, PhD, FRCSEd

- Department of Ophthalmology, Oulu University Hospital, Oulu, Finland.

- Department of Radiology, Oulu University Hospital, Oulu, Finland.

- Institute of Biomaterials, Tampere University of Technology, Tampere, Finland.

- Department of Surgery, Oulu University Hospital, Oulu, Finland.

Purpose: To measure the amount and duration of indentation depth achieved with biodegradable poly-L/D-lactide 96/4 (PLA96) and silicone sponge implants.

Methods: Thirty rabbits underwent a scleral buckling procedure. A PLA96 buckling implant was used in 15 and a silicone sponge buckling implant in 15 rabbits. A circumferential scleral buckling implant was sutured episclerally on the left eye of each rabbit, just temporal to the superior rectus muscle and 7 mm posterior to the limbus. CT scanning was performed at one week, and three and five months postoperatively.

## Results: The PLA96 buckling implant (implant diameter 3–3.5 mm) used in this study created lower indentation than the silicone sponge implant (implant diameter 4 mm), possibly as a result of the surgical technique. The indentation created by the PLA96 implant decreased over time compared with the silicone implant. There were no complications related to either kind of implant.

Conclusion: Both silicone sponge implants and PLA96 implants caused indentation that decreased in a comparable manner over the follow-up period (5 months).

Brief summary statement: An experimental study: bioabsorbable implants made of poly-(L/D)-lactide 96/4 were sutured episclerally in 15 rabbits; in the control group silicone sponge implants were used. The indentation achieved decreased in both groups in a comparable manner over the follow-up period of 5 months. There were no complications related to either kind of material.

Key words: bioabsorbable, experimental, polylactide, retinal detachment, scleral buckling